## Slide 1: THE LHIRES-III SPECTROGRAPH
© C2PU, Observatoire de la Côte d’Azur,
Université de la Côte d’Azur
Jean-Pierre Rivet
CNRS, OCA,
Dept. Lagrange
jean-pierre.rivet@oca.eu
Version 04, 14/11/2017

## Slide 2: The LHIRES-III
LHIRES = Littrow High RESolution spectrograph
The spectrograph
used at C2PU:
The LHIRES-III
Manufacturer:
Sheliak Instruments
http://www.shelyak.com/

## Slide 3: Contents
Basics about spectrographs
Basics about diffraction gratings
The LHIRES-III spectrograph
Practical issues
ESO reference spectra
Sample spectra

## Slide 4: BASICS ABOUT SPECTROGRAPHS

## Slide 5: A wide variety of spectrographs
Slit spectrographs
Fiber(s) spectrographs
Echelle spectrographs
Integral Field Unit spectrographs
Fourier transform spectrographs
Objective prism spectrographs
...
UVES echelle spectrograph on
the UT2 telescope at VLT (ESO).
GIRAFFE spectrograph on
the UT2 telescope at VLT (ESO).
SOPHIE spectrograph on
the 193cm telescope at OHP
ELODIE spectrograph on
the 193cm telescope at OHP

## Slide 6: A wide variety of spectrographs
Several ways to select photons in the focal plane of the telescope:
Slit
Optical fibers
None
Several ways to do chromatic dispersion:
Prism(s)
Grating(s)
Grism (a combination of grating and prism)
...
Several spectral domains:
IR
Visible
UV
...
Several spectral resolutions:
Low resolution (for overall spectrum shape)
High resolution (for line profile sampling or radial velocities)

## Slide 7: Building blocks
“Ordinary” slit spectrograph
The slit and slit environment:
located in the telescope’s focal plane,
selects the region of the field of view to analyze.
The collimation optics:
produces a image of the slit at infinity,
produces an image of the telescope’s pupil on the dispersive element.
The dispersing element:
introduces wavelength-dependent angular shifts.
The camera optics:
forms a sharp image of the slit on the sensor’s sensitive surface
The sensor (nowadays mostly CCD cameras):
records the spectral informations.
The calibration sources:
Tungsten filament bulb for flat-fielding (black body source),
spectral lamps (Th-Ar, Ne-Ar,...) for spectral calibration.

## Slide 8: Building blocks
Dispersing element
(grating)
r
i
Collimated input beam
Collimation optics
Dispersed beam
Camera optics
Sensor
Entrance slit
Light from
the telescope
Ne
Calibration
sources

## Slide 9: Littrow configuration
i
r
Side view
Top view
Littrow condition: r = i
Collimator optics = Camera optics
(cost effective configuration)

## Slide 10: BASICS ABOUT DIFFRACTION GRATINGS

## Slide 11: Diffraction by 1 element
Incident beam
assumed parallel
(wavelength )
Diffracted beam
Collimator
Screen

b
Non-reflecting
substrate
Maximum in the direction of geometric optics: r = - r
Typical angular width: ~  / b
i
r
Reflecting
element
f

## Slide 12: Diffraction by “n” elements
Incident beam
(wavelength )
Screen
?
Non-reflecting
substrate
Reflecting
elements
Collimator
i

## Slide 13: Diffraction by “n” elements
Incident beam
(wavelength )
i
Screen
?
Collimator


## Slide 14: Diffraction by “n” elements
Screen
Collimator
NO LIGHT !
Diffracted beams out of phase :
destructive interferences
 NO LIGHT

## Slide 15: Diffraction by “n” elements
Screen
Collimator
LIGHT !
Diffracted beams in phase :
constructive interferences
 MAXIMUM LIGHT

## Slide 16: Diffraction by “n” elements
a
a
a
a
i
Ray 0
Ray 1
Delay of Ray 1 wrt Ray 0 = a sin(i)

## Slide 17: Diffraction by “n” elements
a
a
a
a
r
Ray 0
Ray 1
Delay of Ray 1 wrt Ray 0 = a sin(r)

## Slide 18: Diffraction by “n” elements
a
a
a
a
r
Ray 0
Ray 1
i
Ray 0
Ray 1
Total delay of Ray 1 wrt Ray 0 :
 = a sin(i) + a sin(r)
Condition for
constructive interferences:
 = k . 
integer; called the “order”

## Slide 19: Diffraction by “n” elements
a
a
a
a
r
Ray 0 ’
Ray 1 ’
i
Ray 0
Ray 1
Order k = 0
Condition for
constructive interferences:
 = 0, whatever 
sin(i) + sin(r) = 0
Snell’s law !
direction of reflection on the grating’s plane
according to geometric optics
NON DISPERSIVE

## Slide 20: Diffraction by “n” elements
a
a
a
a
r
Ray 0 ’
Ray 1 ’
i
Ray 0
Ray 1
Order k ≠ 0
Condition for
constructive interferences:
 = k . 
sin(i) + sin(r) = k .  / a
DISPERSIVE

## Slide 21: Focal pattern
r
i
Fraunhoffer’s diffraction
integral leads to:
with
Grating function
G()
Diffraction function
D()

## Slide 22: Grating function
Grating function
G()
Diffraction function
D()

G()
N2
0
+1
+2
-1
-2
order -2
order -1
order 0
order +1
order +2

## Slide 23: Diffraction function
Grating function
G()
Diffraction function
D()

D()
0
+1
+2
-1
-2
0.047
1
0.016
0.047
0.016

## Slide 24: Diffraction pattern (monochr.)
sin(i) + sin(r)
0
Relative intensity
~  / (N.a)
~  / a
~  / b
Diffraction envelope
 / a
2 / a
3 / a
- / a
-2 / a
-3 / a
Thickness of the
diffraction lines
Order
separation
Width of the
diffraction
envelope
a
a
a
a
N
b
N.a

## Slide 25: Scale correspondance
3 length scales for the grating:
Total width: N.a
Pitch: a
Diffractive elements width: b
3 length scales for diffracted pattern:
Thickness of the line:  / N.a
Order separation:  / a
Width of the diffraction envelope:  / b

## Slide 26: Diffraction pattern (polychr.)
sin(i) + sin(r)
0
Relative intensity
 / a
2 / a
3 / a
- / a
-2 / a
-3 / a
Order 0:
non dispersive
Order 1:
dispersive
Order 2:
more dispersive
Order 3:
even more dispersive

## Slide 27: Blazed gratings
r
i
: Normal to the grating surface
: Normal to the diffracting surfaces
 (blaze angle)
i
r
STANDARD GRATING
BLAZED GRATING

## Slide 28: Focal pattern
Fraunhoffer’s diffraction
integral leads to:
with
Grating function
G()
Diffraction function
D()
: Normal to the grating surface
: Normal to the diffracting surfaces
 (blaze angle)
i
r
i
r

r
i

## Slide 29: Blazed vs standard gratings
r
i
Diffraction envelope is maximum when:
r = - i
0th order is maximum when:
r = - i
Diffraction envelope is maximum when:
r = - i
0th order is maximum when:
r = - i
: Normal to the grating surface
: Normal to the diffracting surfaces
 (blaze angle)
i
r
i
r
STANDARD GRATING
BLAZED GRATING

( i = i -  , r = r -  )

## Slide 30: Diffraction pattern (polychr.)
sin(i) + sin(r)
0
Relative intensity
 / a
2 / a
3 / a
- / a
-2 / a
-3 / a
Order 0:
non dispersive
Order 1:
dispersive
Order 2:
more dispersive
Order 3:
even more dispersive
STANDARD GRATING

## Slide 31: Diffraction pattern
sin(i) + sin(r)
0
Relative intensity
 / a
2 / a
3 / a
- / a
-2 / a
-3 / a
Maximum of
diffraction curve
on order k ≠ 0
BLAZED GRATING
Blaze angle  depends on
the central wavelength 0
and order k

## Slide 32: Real blazed gratings
“Real” gratings are
more like that:
Carved in a piece of glass by a very hard (diamond) tool with wanted profile and angle, driven by a machine with a high precision master screw translator, for highly regular groove spacing.

## Slide 33: Spectral resolution
A single, infinitely sharp emission line...

I()
0

I()
0
... will lead to:
Convolution with
instrumental response

Spectral resolution: R =

## Slide 34: Spectral resolution

I()
0

I()
0
Convolution with
instrumental response


I()
0

I()
0
Convolution with
instrumental response

1

1
1 - 0 > 
1 - 0 < 
Unresolved
Resolved
Spectral resolution = smallest visible wavelength span 
divided by central wavelength 0

## Slide 35: Spectral resolution

I()
0
Instrumental response:

= convolution of grating function with image of slit
Spectral resolution depends
on the grating and on the slit.
Can also be limited by CCD’s
pixel size...

## Slide 36: THE LHIRES-III SPECTROGRAPH

## Slide 37: The LHIRES-III

## Slide 38: The LHIRES-III
Micrometric screw
(to tilt the gating)
Diffraction
blazed grating)
Collimator / camera
optics
Bending mirror
Science camera
Guiding camera
Focuser for the
guiding camera
Slit
environment
Bending mirror
F/12.5 input beam
from the telescope

## Slide 39: The LHIRES-III
Micrometric screw
(to tilt the gating)
Diffraction
blazed grating)
Collimator / camera
optics
Bending mirror
Science port
Guiding port
Focuser for the
guiding camera
Slit
environment
Bending mirror
F/12.5 input port

## Slide 40: The LHIRES-III

## Slide 41: The slit environment
Bending flat mirror
Input beam
(from telescope)
Guiding output port
Input slit
Output port
focusing optics
Slit environment

## Slide 42: The slit environment
15 m slit
19 m slit
Active slit
25 m slit
35 m slit
Optically polished component:
MUST HE HANDELED WITH CARE
Slit holding plate

## Slide 43: The slit environment
35
25
15
19
Hole
Slit holding plate
Active slit:
C1

## Slide 44: The calibration lamps
Spectral calibration lamp:
a small glass bulb filled with low pressure gases, producing
strong and well-defined emission lines when an
electric current passes through. Example Neon-Argon.
Goal: determine the pixel-wavelength relationship.
Flat calibration lamp:
A Tungsten filament bulb producing a “black body”
continuous spectrum.
Goal: determine the photometric throughput variations,
pixel by pixel.
Spectral calibration switch
Flat calibration switch
Power supply plug

## Slide 45: The Neon-Argon spectrum

## Slide 46: The diffraction gratings
Protection frame
Active grating surface
Housing
Tilt axis
High precision optical component:
MUST HE HANDELED WITH EXTREME CARE
NO FINGER PRINTS !
Available gratings:
150 gr/mm
300 gr/mm
2400 gr/mm

## Slide 47: Spectral resolution
Available slits:
15 microns
19 microns
25 microns
35 microns
 | 15 m | 19 m | 25 m | 35 m
150 gr/mm | 890 | 750 | 610 | 470
300 gr/mm | 1780 | 1510 | 1230 | 940
2400 gr/mm | 23000 | 19500 | 15900 | 12100
Slit
Grating
Spectral resolution @ 656.3nm (H)
Available gratings:
150 gr/mm
300 gr/mm
2400 gr/mm
Conditions:
Telescope aperture ratio F/D = 9.0

## Slide 48: Spectral range, spectral scale
Spectral range accessible on a single image, and spectral scale
@ 656.3nm (H)
Grating | Spectral range | Spectral scale
150 gr/mm | 230 nm | 0.30 nm/pix
300 gr/mm | 114 nm | 0.15 nm/pix
2400 gr/mm | 9 nm | 0.012 nm/pix
Conditions:
Science camera: SBIG ST402 (pix=9m, Size=6.9mm)

## Slide 49: PRACTICAL ISSUES

## Slide 50: Preliminary adjustments
Operations to be done before the LHIRES-III is set on the telescope:
Guiding camera focusing.
Guiding camera orientation adjustment.
Science camera orientation adjustment.
Science camera focusing.
Micrometric screw pre-calibration.

## Slide 51: Guiding camera focusing
Guiding optics should produce a
sharp image of the slit onto the guiding camera’s sensor.
Unfocused slit image
on the guiding camera
Focused slit image
on the guiding camera

## Slide 52: Guiding camera focusing
iNova
guiding
camera
Focus locking
scroll wheel
Telescope
interface (2”)
LHIRES-III
Guiding camera
mount
Guiding camera
lock screws (x3)
Switch off all calibration lamps.
Use an external light source to light the slit through the telescope interface tube (2”).
Untighten the focus locking scroll wheel.
Untighten the 3 guiding camera lock screws.
Rotate the guiding camera mount while maintaining the guiding camera in an approximately fixed orientation, until the slit image on the guiding camera display appears sharp.
Tighten the focus locking scroll wheel.
Check that the slit image is still sharp.
Tighten the 3 guiding camera lock screws.
Light
source

## Slide 53: Guiding camera orientation
iNova
guiding
camera
ST402
science
camera
Telescope
interface (2”)
LHIRES-III
Camera orientation
angle adjustment
Guiding
port
Science
port

## Slide 54: Guiding camera orientation
Guiding optics should produce a
sharp image of the slit onto the guiding camera’s sensor.
Wrong orientation of the slit image
on the guiding camera
Correct orientation of the slit image
on the guiding camera

## Slide 55: Guiding camera orientation
iNova
guiding
camera
Focus locking
scroll wheel
Telescope
interface (2”)
LHIRES-III
Guiding camera
mount
Guiding camera
lock screws (x3)
Switch off all calibration lamps.
Use an external light source to light the slit through the telescope interface tube (2”).
Untighten the 3 guiding camera lock screws.
Rotate the guiding camera in its mount, until the slit image on the guiding camera display appears parallel to the reticule’s horizontal line.
Tighten the 3 guiding camera lock screws.
Light
source

## Slide 56: Science camera orientation
iNova
guiding
camera
ST402
science
camera
Telescope
interface (2”)
LHIRES-III
Camera orientation
angle adjustment
Guiding
port
Science
port

## Slide 57: Spectral image geometry
Typical non-ideal spectral image
Sky background
Absorption
lines
Continuum
Emission
line

## Slide 58: Spectral image geometry
Camera pixel rows orientation
Dispersion direction
Spectral lines direction
slant angle
tilt angle
Spectral lines curvature:
“smile”
Tilt: dispersion vs pixel rows. Depends on camera’s orientation w.r.t. grating.
Slant: spectral lines vs pixels columns. Depends on camera’s orientation w.r.t. slit.
Smile: spectral lines curvature. Depends on inner optical aberrations.

## Slide 59: Spectral image geometry
Ideal spectral image
Tilt: = 0.
Slant: = 0.
Smile: = 0.

## Slide 60: Science camera orientation
iNova
guiding
camera
ST402
science
camera
Camera orientation
angle adjustment
Guiding
port
Science
port
Switch the Ne-Ar calibration lamp on.
Adjust the camera orientation angle should be set so that slant angle be 0
Tighten gently the camera positioning screws.
Tilt angle and smile will never be 0. They depend on the internal structure of the LHIRES-III.  Will be software-corrected.

## Slide 61: Science camera focusing
Unfocused spectral image
of the Ne-Ar calibration lamp
Focused spectral image
of the Ne-Ar calibration lamp
Littrow (collimation+camera) optics should produce a
sharp image of the slit onto the science camera’s sensor.
DON’T GET CONFUSED WITH THE TELESCOPE’S FOCUSING :
(tuning the telescope’s focus position so that stellar images
are sharply formed on the slit’s surface).

## Slide 62: Science camera focusing
iNova
guiding
camera
ST402
science
camera
Screw
Lateral flap
Handles
(do not unscrew)
Switch the Ne-Ar calibration lamp on.
Remove the locking screw of the side flap.
Grab the handles of the flap and slide it so as to unlock it.
Remove the side flap.
Do the same on both sides.

## Slide 63: Science camera focusing
iNova
guiding
camera
ST402
science
camera
Unlock nylon screws on both sides.
Adjust the focus scroll wheel to have sharp spectral lines on the camera display.
Lock the nylon screws on both sides.
Check the line’s sharpness.
Close the flaps on both sides.
Littrow optics
barrel
Nylon locking screw
Focus scroll wheel

## Slide 64: The micrometric screw to adjust the grating angle
Fixed tilt axis
Active grating surface
0
5
10
15
20
25
30
35
40
45
Last visible
mark: 23.5
Drum tick mark
in front of the
fixed index : 34
How to read the
micrometric screw :
Fixed index
Value = 23.5+0.34 = 23.84
Micrometric screw
Half-integer
tick marks
Integer
tick marks
Useful to select the
central wavelength !

## Slide 65: Micrometric screw calibration
Micrometric screw
Chart: Calibration of the micrometric screw
| 4764 | 4879 | 5017 | 5187 | 5495 | 5606 | 5852.49 | 5881.89 | 5944.83 | 5975.53 | 6030 | 6074.34 | 6096.16 | 6143.06 | 6163.59 | 6217.28 | 6266.49 | 6304.79 | 6334.43 | 6382.99 | 6402.25 | 6506.53 | 6532.88 | 6598.95 | 6678.28 | 6717.04 | 6929.47 | 7032.41
 | 2.9159 | 2.9317 | 2.9508 | 2.9742 | 3.0167 | 3.032 | 3.0659 | 3.07 | 3.0787 | 3.0829 | 3.0904 | 3.0965 | 3.0995 | 3.106 | 3.1088 | 3.1162 | 3.123 | 3.1283 | 3.1324 | 3.1391 | 3.1417 | 3.1561 | 3.1597 | 3.1688 | 3.1798 | 3.1851 | 3.2144 | 3.2286
 | 2.91 | 2.94 | 2.96 | 2.98 | 3.01 | 3.02 | 3.055 | 3.06 | 3.075 | 3.08 | 3.09 | 3.1 | 3.105 | 3.11 | 3.112 | 3.12 | 3.129 | 3.132 | 3.136 | 3.141 | 3.1415 | 3.153 | 3.158 | 3.169 | 3.178 | 3.185 | 3.214 | 3.23
4500
5000
5500
7000
6500
6000
7500
Wavelength (Å)
3.25
3.20
3.15
3.10
3.05
3.00
2.95
2.90
2.85
Micrometer (mm)
For further on-sky operations, we need a coarse relationship between the micrometer reading and the central wavelength on the science spectroscopic image.

## Slide 66: Micrometric screw calibration
Switch the Ne-Ar calibration lamp on.
Screw the micrometer almost entirely, until a very bright single line appears.
Put this line (zeroth order) in the middle of the science camera image, and read the micrometer’s position.

## Slide 67: Micrometric screw calibration
Unscrew the micrometer until several spectral lines appear (first order).
Try to recognize a given spectral line (e.g. the 6344.43Å line) from the calibration lamp’s spectral atlas (available upon request).
Put this line in the middle of the science camera image, and read the micrometer’s position. Store the couples [wavelength-micrometer reading] in a spreadsheet (template available upon request).
6334.43Å line
6334.43Å line

## Slide 68: Micrometric screw calibration
Do it again for as much lines as possible.
Compute the best fitting polynomial by a least square method. A first order polynomial is often sufficient.
This pre-calibration relationship will serve during on-sky observing sessions, to find the micrometer position to choose, in order to have the wanted spectral region in the center of the spectral image.
Chart: Calibration of the micrometric screw
| 4764 | 4879 | 5017 | 5187 | 5495 | 5606 | 5852.49 | 5881.89 | 5944.83 | 5975.53 | 6030 | 6074.34 | 6096.16 | 6143.06 | 6163.59 | 6217.28 | 6266.49 | 6304.79 | 6334.43 | 6382.99 | 6402.25 | 6506.53 | 6532.88 | 6598.95 | 6678.28 | 6717.04 | 6929.47 | 7032.41
 | 2.9159 | 2.9317 | 2.9508 | 2.9742 | 3.0167 | 3.032 | 3.0659 | 3.07 | 3.0787 | 3.0829 | 3.0904 | 3.0965 | 3.0995 | 3.106 | 3.1088 | 3.1162 | 3.123 | 3.1283 | 3.1324 | 3.1391 | 3.1417 | 3.1561 | 3.1597 | 3.1688 | 3.1798 | 3.1851 | 3.2144 | 3.2286
 | 2.91 | 2.94 | 2.96 | 2.98 | 3.01 | 3.02 | 3.055 | 3.06 | 3.075 | 3.08 | 3.09 | 3.1 | 3.105 | 3.11 | 3.112 | 3.12 | 3.129 | 3.132 | 3.136 | 3.141 | 3.1415 | 3.153 | 3.158 | 3.169 | 3.178 | 3.185 | 3.214 | 3.23
4500
5000
5500
7000
6500
6000
7500
Wavelength (Å)
3.25
3.20
3.15
3.10
3.05
3.00
2.95
2.90
2.85
Micrometer (mm)

## Slide 69: Typical observing session
Point the telescope to the first science target.
Switch the Tungsten calibration lamp on, and record 5 flat frames.
Switch the Tungsten calibration lamp off. Switch the Ne-Ar calibration lamp on, and record 1 spectral calibration frame. Switch the Ne-Ar calibration lamp off.
Record science frames for at most 15 minutes.
Record 5 flat frames and 1 spectral calibration frame (see items 2 and 3).
Point the telescope to a spectral reference star (e.g. Vega; list available upon request), as close as possible to the science target.
Record 5 science frames on the reference star (as in item 4).
Loop back to item 2, or to item 1 for a new science target.

## Slide 70: Observing log
LHIRES-III OBSERVING LOG
Date : 2017/02/15
Observers : Jean DUPONT, Michel DUPOND.
Telescope : Epsilon@C2PU (OCA)
Instrument : LHIRES-IIIa
Science camera: SBIG ST402 FW
Guiding camera: iNova PLB-Mx
OBSERVING RUN NUMBER 01 ---------------------------------------
.
.
---------------------------------------------------------------
OBSERVING RUN NUMBER 02 ---------------------------------------
.
.
---------------------------------------------------------------
FILE HEADER:

## Slide 71: Observing log
OBSERVING RUN NUMBER 01 ---------------------------------------
Target : Jupiter
Reference star : Vega
Grating : 150 tr/mm
Micrometer : 03.16mm
Central wavelength : 656.3nm
Telescope on target: 22:30:00 UT
22:30:30 UT
05 Tungsten flat frames. Exp=0001s000
22:30:50 UT
05 Argon-Neon calibration frames. Exp=0000s100
22:31:00 UT
10 science frames on Ref. star Exp=0000s500
22:32:00 UT
05 Tungsten flat frames. Exp=0001s000
22:32:30 UT
05 Argon-Neon calibration frames. Exp=0000s100
22:33:10 UT
10 science frames on target. Exp=0000s500
.
.
.
22:35:50 UT
---------------------------------------------------------------
OBSERVING RUNS:

## Slide 72: European South Observatory REFERENCE SPECTRA

## Slide 73: ESO reference spectra
https://www.eso.org/sci/observing/tools/standards/spectra/stanlis.html
Right ascension-ordered list of standards:
Click the name
of the selected star

## Slide 74: ESO reference spectra
Basic informations
on the star:
Finder
Magnitudes, spectral types
Coordinates
AB mag. spectrum
Flux spectrum
Go to FTP server to download

## Slide 75: ESO reference spectra
ftp://ftp.eso.org/pub/stecf/standards/ctiostan/
FTP for downloading standard spectra:
Click the name
of the selected star

## Slide 76: ESO reference spectra
Spectral data:
Wavelength (Å)
Flux (ergs/cm2/s/Å x 1016)
f-prefixed files
(ex: “fhr9087.dat”):
Flux (milli-Jansky)
Bin size (Å)
1 Jansky = 10-26 W m-2 Hz-1

## Slide 77: ESO reference spectra
Spectral data:
Wavelength (Å)
AB magnitude
m-prefixed files
(ex: “mhr9087.dat”):
Bin size (Å)
mAB=-2.5×Log10( Flux / 3631Jy )

## Slide 78: SAMPLE SPECTRA

## Slide 79: Solar spectrum
The Hydrogen “H” line (LHIRES-III + 2400 gr/mm)

## Slide 80: Solar spectrum
The Sodium D1 and D2 lines (LHIRES-III + 2400 gr/mm)

## Slide 81: Solar spectrum
The Magnesium triplet (LHIRES-III + 2400 gr/mm)

## Slide 82: The rotation of saturn
The Hydrogen H line in Saturn’s spectrum (LHIRES-III + 2400 gr/mm)
The lines are tilted by the planet’s surface rotation (Doppler effect)

## Slide 83: END